## How to switch between trapping and no-trapping mode

- You can look at the home page on the EASY to see which mode it currently is in

Without trap shows a straight line 				With trap shows the trap column

- To switch the trapping mode click on the Maintenance/Devices/EAY-nLC [HPLC]/Properties

- To select no-trap mode check the box next to “One column setup”

- To select trapping mode uncheck the next to “One column setup”

- Hit “Apply” and accept the popup confirmation window

No-trap mode (using analytical column only) 		     Trap mode (using trap and analytical column)

### Make sure your pluming is setup correctly

- In trap mode put a trap column in front of the analytical column

- In no-trap mode replace the trap column with an empty fused silica

- That’s it.

### Now make sure the sample loading in the method is setup correctly

- For no-trap mode we recommend to load at lower flowrates e.g.  0.3 µl/min, or set the max pressure instead e.g. 150-200 bar

- the loading volume should be: 2 x sample volume + 2µl

- But always use a minimum of 6 µl (1-2 µl of sample volume)

- e.g. if you load 4 µl of sample: 2 x 4 µl + 2 µl = 10 µl

- For trap mode you can load at higher flowrates e.g. 2.5 µl, or set the max pressure instead e.g. 150-200 bar

- the loading volume should be: 2 x sample volume + 2µl

- But always use a minimum of 6 µl (1-2 µl of sample volume)

- e.g. if you load 4 µl of sample: 2 x 4 µl + 2 µl = 10 µl